[image: image1.png]



[image: image2.emf]







Faculty of Science
 Appendix XII – Laser Laboratory Entry Requirements

· All laser laboratory entry doors must be fitted with the correct laser warning signs and labels that comply with AS/NZS IEC 60825.14:2011 and Appendix III.

· All laser laboratory entry doors must also be fitted with a red warning light of diameter no less than 5 cm. This light must be visible and switched on when the laser is in operation.

· All laser laboratory entry doors must have a sign (eg. whiteboard) indicating what laser/s is/are operating and the PPE requirements that are needed in order to enter.

· All laser laboratory entry doors must be locked at all times.

· Where possible, all laboratories/rooms in which a Class 3B or Class 4 laser is operated must be fitted with a door interlock that will isolate the laser hazard if the door is opened. The operator in charge of the laser is permitted to momentarily override the interlock to allow entry to other authorised persons if it safe to do so.

· Interlocked doors will consist of a fail open door strike mechanism, an entry keypad, an entry panic/stop button and an exit interlock override button. These components plus the magnetically actuated non-contact safety interlock must be switched on from inside the laser laboratory by the laser operator when the laser is in use.
· All laser laboratory entry doors must be fitted with a buzzer/door bell that alerts the laser operator when activated. This buzzer must be activated by personnel wishing to gain access to a laser laboratory.
Regardless of the level of training, experience and/or familiarity you have with any of the laser labs, entry to a laser lab can only be in accordance with the lab’s door sign. If the door sign indicates the laser is switched on, you must wait to be let in by the laser operator. If no one is present in the laser lab, you may enter the lab (although you are not advised to) provided you adhere to the safety requirements outlined on the lab door.

Upon gaining entry to a laser lab, any hazards that you encounter (meaning the lab was not left in a safe state) must be removed immediately and reported to the appropriate lab supervisor. NB: This may include switching off the laser system.
2nd September 2015

Page 1 of 1

